INVIZE AB

EQUIPMENT VERIFICATION PROTOCOL

| Device Under Test: | Optical Flicker Generator | Date of Test: | 2026-05-18 |
| --- | --- | --- | --- |
| Model / Revision: | Rev 1 | Location: | ESS, Lund, Sweden |
| Serial Number: | prototype-1 | Operator: | Andreas Hagelberg |
| Ambient Temp: | 20 °C | Relative Humidity: | 30% RH |

1. Reference System Conditions

Measurements were captured using an Aim-TTi TF930 Universal Frequency Counter with an internal 10 MHz TCXO base clock source. Standard baseline specification limits determine the combined reference standard uncertainty profile to be ±1.2 ppm. Raw readout digits were manually truncated from the faceplate to address fluctuating sub-millihertz jitter.

2. Verification Data Table

| Nominal / Set Frequency (Hz) | Measured Frequency (Hz) | Absolute Error (Hz) | Relative Error (ppm) | TF930 Timebase Uncertainty (Hz) |
| --- | --- | --- | --- | --- |
| 1 | 1.0001 | +0.0001 | +100.00 ppm | ±0.0000012 |
| 10 | 10.001 | +0.0010 | +100.00 ppm | ±0.0000120 |
| 50 | 50.006 | +0.0060 | +120.00 ppm | ±0.0000600 |
| 100 | 100.01 | +0.0100 | +100.00 ppm | ±0.0001200 |
| 200 | 200.02 | +0.0200 | +100.00 ppm | ±0.0002400 |
| 500 | 500.06 | +0.0600 | +120.00 ppm | ±0.0006000 |

3. Characterization Summary

The Optical Flicker Generator demonstrates a predictable, highly linear systematic offset grouped consistently between +100.00 ppm and +120.00 ppm across its test envelope. The uniform distribution suggests a stable clock scaling profile in the base timer/oscillator subsystem rather than dynamic firmware lag. Reference instrument uncertainty remains two orders of magnitude below observed deviations, confirming statistical validity of the performance window.

4. Approvals and Sign-off

______________________________
Andreas Hagelberg
Invize AB
Date: 2026-05-18